| Force | Pressure | Temperature | Switch

**Pressure sensors for industrial applications Model P3297**


### Non linearity 0.5% (option 0.25%)

**Standard output: 4...20 mA; 2-wire**

**or 0...5 VDC; 3-wire or 0...10 VDC; 3-wire or 0.5 4.5 VDC; 3-wire**

**or 0.5 4.5 VDC ratiometric**

**Description**

Robustness and long-term stability during operation are the strengths of this compact pressure sensor for general industrial applications.

The materials and technologies used make these sensors suitable for applications with aggerssive media. Welded connections between pressure cell and process connection require no sealing elements and make the measuring system particularly resistant to mechanical shock and vibration. The compact design makes these sensors interesting for room critical applications.

A wide variety of electrical connections and pressure ports simplifies the adaptation to different applications. The pressure sensor is internationally certified and ready for global deployment.

The pressure sensors comply with electromagnetic compatibility requirements (EMC) as per EN 61326.

# Features

O Measuring range from 0...1 bar to 0...600 bar O Medium wetted parts of stainless steel

O High EMC protection according to EN 61 326 O Compact instrument size

O No internal sealing elements

O Highly resistance to shock and vibration O For dynamic or static measurements

# Measuring range

Gauge pressure 0…1 bar to 0…600 bar

-1…0 bar to -1…+24 bar

# Applications

Hydraulics and pneumatics Pumps and compressors Building automation

Test stand construction

Machine and apparatus construction

**Model: P3297**



|  |  |  |  |  |
| --- | --- | --- | --- | --- |
| tecsis GmbHCarl-Legien Str. 40-44D-63073 Offenbach / Main | Sales National | Sales International | e-Mail: info@tecsis.de | DE **7**08 k 02/2016 |
| Tel.: +49(0) 69 / 5806-0 | Fax: +49(0) 69 / 5806-7788 | Fax: +49(0) 69 / 5806-7788 | Internet: [www.tecsis.de](http://www.tecsis.de/) | p. **1** / 4 |

# Technical Data

|  |  |
| --- | --- |
| **Model** | **P3297** |
| Pressure type | positive and negative gauge pressure absolut pressure on request |
| - Measuring range [bar] | 0...1 bar to 0…600 bar-1..0 bar to -1…+24 bar |
| - overrange limit [bar] | x 2 |
| - burst pressure [bar] | x 6 |
| Sensor element | piezoresistive to 0..6 bar, thin film as of 0..10 bar |
| Output signal | 4...20 mA0...5 VDC1...5 VDC0...10 VDC0.5 4.5 VDC0.5 4.5 VDC | 1. wire
2. wire

3- wire3- wire3- wire ratiometric |
| Non linearity1) |  0.5% of F.S.; option: 0.25% of F.S. |
| Accuracy 2) |  1.0% of F.S.; option: 0.5% of F.S. 3) |
| Zero Drift |  0,5% of F.S. (typ.),  0,8% of F.S. (max.) |
| Non-repeatability |  0.1% of F.S. |
| Long-term Drift |  0.1% of F.S. (by reference conditions) |
| Materialnot wetted parts medium wetted parts | stainless steel 316L, HNBR, PAstainless steel 316L (from 0…10 bar rel. 13-8PH) |
| Pressure connection | G 1/4 according to DIN 3852-E G 1/4 according to EN 837G 1/2 according to EN 837 1/4 NPT1/2 NPT, other pressure connection on request |
| Electrical connection | connector DIN EN 175301-803 Form A with junction box connector DIN EN 175301-803 Form C with junction box circular plug-in connector M12x1 (4-pin)cable outlet: 2m, other electrical connection on request |
| Power supply / load 4...20 mA | 8...30 VDC | RA []  (UB [V] – 8V) / 0.02A |
| 0…1...5 V | 8...30 VDC | > max. output / 1 mA |
| 0...10 V | 14...30 VDC | > max. output / 1 mA |
| 0.5 V | 8…30 VDC | > max. output / 1 mA |
| 0.5 ... 4.5 V ratiometric | 5 VDC ± 10% | RA > 4.5k |
| Reponse time |  4ms within 10% to 90% of F.S. |
| RoHS-conformity | yes |
| Approval according to | cULus |
| CE-conformance | 2004/108/EWG interference emission and interference resistance to EN 61 326 interference emission limit class B97/23/EG pressure gauge code |
| Electrical protections | polarity, overvoltage and short-circuit protection |
| Ingress protection(per IEC 60529)4 | Plug DIN EN 175301-803: IP 65Circular Connectors M12x1: IP 67 Cable output: IP 67 |
| Temperature influence |  1% typ.  2.5% max.in range 0…80°C |
| Temperature ranges |  |
| compansated range | 0...80°C |
| storage | -20...80°C (Option: -30..100°C) |
| media | 0...80°C (Option: -30..100°C) |
| ambient | 0...80°C (Option: -30..100°C) |
| Load capacityshock (mechanical) vibration (under resonance) | 500g acc. to IEC 60068-2-2710g acc. to IEC 60068-2-6 |
| Weight | approx. 80g |

protection.

1) According to IEC 61298-2

2) Including non linearity, hysteresis, non repeatability, variation of zero point and finale value (is equal to error according to IEC 61298-2).

3) By option: accuracy 0.5% and signal 0…5V is accuracy 0.6%

4) The specified protection class (according to IEC 60529) only applies when plugged in using mating connectors with corresponding

# Dimension (mm)

### Case

connector according to DIN EN 175301 – 803 Form A

connector according to DIN EN 175301 – 803 Form C

circular plug-in connector M12x1

Cable outlet



  

A

clamping range

clamping range

### Pressure connections

G 1/2 B G 1/4 B G 1/8 B G 1/2 DIN 3852-E

G 1/4 A DIN 3852-E 1/2 NPT 1/4 NPT R 1/2

R 3/8 R 1/4 G 1/4 female 1/4 NPT female



M20 x 1,5



# Electrical connector

### Two-wire system

Connector according to DIN EN 175301-803 Form A with junction box

Connector according to DIN EN 175301-803 Form C with junction box

UB / S+

1

3

2

0V / S-

UB / S+

1

3

2

0V / S-

E-001

E-005

Circular plug-in connector M12x1 Cable outlet

braun

UB / S+

4 3

1 2

0V / S-

UB / S+

brown

0V / S-

blue

E-033

E-015

blau

### Three-wire system

Connector according to DIN EN 175301-803 Form A with junction box

Connector according to DIN EN 175301-803 Form C with junction box

##

UB

S+

0V / S-

UB

S+

 0V / S-

E-002 E-006

Circular plug-in connector M12x1 Cable outlet

braun

UB

S+

4 3

1 2

0V / S-

UB

brown

S+ black

 0V

blue

E-034

E-017

/ S- grün

schwarz

blau

Modifications reserved